| α/α | Περιγραφή | TEXNIKEΣ ΠΡΟΔΙΑΓΡΑΦΕΣ | ΠΡΟΥΠΟΛΟΓΙΣΜΟΣ |
| --- | --- | --- | --- |
| 1 | Αναβάθμιση μαγνητομέτρων: Physical Properties Measuring System 9 Tesla (PPMS) της εταιρίας QUANTUM Design, SQUID magnetometer 5.5 Tesla της εταιρίας Quantum Design | - Μηχανική Αντλία κενού δύο σταδίων με τις ακόλουθες προδιαγραφές: Α) Flange in/out : DN16 ISO KF Β) Ultimate Pressure : 3 .10-3mbar Γ) Ταχύτητα άντλησης στα 50Hz, μεγαλύτερη ή ίση από 5m3/h Δ) Παγίδα εισόδου αντλίας κενού 0.5 lt (δοχείο σφαιριδίων αλούμινας) Ε) 20 lt λάδι αντλίας και 1 lt σφαιρίδια αλούμινας | 3.000,00 Ευρώ |
|  |  | - Lockin amplifier με τις ακόλουθες προδιαγραφές: Α) Signal Channel - Voltage Inputs Single-ended (A) or differential (A−B) - Sensitivity (Output Scale) 1 nV to 1 V (voltage input) - 1 fA to 1 μA (current input) - Input Impedance 10 MΩ+25 pF, ac (>1 Hz) or dc coupled - Input Range 10 mV to 1 V (peak); max input before overload - Gain Accuracy 1% below 200 kHz and 2% to 4 MHz (signal amplitude <30% of input range) - Input Noise 2.5 nV/√Hz above 1 kHz, 10 mV input range (typical) - CMRR Greater than 90 dB at 1 kHz (dc Coupled, 10 mV to 100 mV input range) - Harmonic Distortion −80 dB below 100 kHz, −60 dB above 100 kHz - Dynamic Reserve Greater than 120 dB - Current Input Ranges 1 μA or 10 nA - Reference Channel - Frequency Range 1 mHz to 4 MHz - Ext Reference Input Impedance 1 MΩ or 50 Ω - Acquisition Time (2 cycles + 5 ms) or 40 ms, whichever is greater - Phase Setting Resolution 360/232 deg - Internal reference: <0.0001° rms at 1 kHz (100 ms, 12 dB/oct) - Time Constants 1 μs to 30 ks - Low Pass Filters Typical RC type filters or Advanced Gaussian/Linear Phase filters - Filter Slope 6, 12, 18, 24 dB/oct rolloffs Β) Internal Oscillator - Frequency 1 mHz to 4 MHz - Frequency Accuracy 25 ppm + 30 μHz with internal timebase - External Timebase 10 MHz timebase input/output on rear panel - Frequency Resolution 6 digits or 0.1 mHz, whichever is greater Γ) Data - Data Channels 4 data channels are displayed - Data Sources Each data channel can be assigned any of these data sources: X, Y, R, θ, Aux In 1–4, Aux Out 1–2, Xnoise, Ynoise, Sine Out Amplitude, Sine Out DC Level, reference phase, fint or fext. Δ) Inputs and Outputs - CH 1 Output Proportional to X or R, ±10 V full scale thru 50 Ω - CH 2 Output Proportional to Y or θ, ±10 V full scale thru 50 Ω - X and Y Outputs Proportional to X and Y, ±10 V full scale thru 50 Ω, rear panel - Aux Outputs 4 BNC D/A outputs, ±10.5 V thru 50 Ω, 1 mV resolution - Aux Inputs 4 BNC A/D inputs, ±10.5 V, 1 mV resolution, 1 MΩ input Ε) Interfaces - IEEE488, RS-232, USB device | 12.000,00 Ευρώ |
|  |  | - Vibrator and its power amplifier με τις ακόλουθες προδιαγραφές - Amplifier to drive the vibrator. - Sine peak force >=16 N - Velocity sine peak >1.4 m/s - Acceleration sine peak >800 m/s2 - Displacement peak to peak <6 mm - Mass of moving element >15 g - G) Usable frequency range dc –to 10 kHz | 4.000,00 Ευρώ |
| 2 | Αναβάθμιση περιθλασίμετρου ακτίνων x. Περιθλασίμετρο SmartLab της εταιρίας Rigaku | 4. Δειγματολήπτης 10 θέσεων Περιθλασίμετρου SmartLab της εταιρίας Rigaku ASC-10 HS Code: 90229000 | 16.000,00 € |
| 3 | Αναβάθμιση φασματόμετρου ΕPR, ER-200D της εταιρείας Bruker | 5. One two-stage Rotary vane pump up to 36 m3/h - Dual-stage, high-performance rotary vane pump with a pumping speed up to 36 m3/h - With 3-phase motor - With temperature sensor (3TF) for optimum motor protection - Integrated gas ballast and HV safety valve - Special feature: Various motor voltages for worldwide employment - Other motor voltages available on request - oil mist filter | 5.000,00 € |
|  |  | 6. Three two-stage Rotary vane pumps up to 12 m3/h - 2 Dual-stage, high-performance rotary vane pump with a pumping speed of up to 12 m3/h - With Single-phase motor - Integrated gas ballast and HV safety valve - Magnetically coupled - oil mist filter | 8.000,00 € |
|  |  | 7. One Liquid helium transfer line - Liquid helium transfer tube for horizontal feed Oxford instruments continuous flow cryostats with flexible section length 1.3 m and dewar leg with length 1m. | 5.000,00 € |